| Device Under Test Model: ML600 DUAL,SYRINGE PUMP ASSY Part #: 61502-01 Serial #: ML600AK1203 Description: SYRINGE PUMP ASSEMBLY ML600 DUAL | Calibration Device Name: Mitutoyo 543-564A Model: ID-H0560 Serial #: 10009710 N.I.S.T #: 821/487256-08 Accuracy: +/-0.0025mm Calibration Due: 04/30/2014 |
| --- | --- |

| Accuracy Specifications: +/- 1% Of Reading @ ≥ 30% of Full Stroke Dispense (60mm) +/- 1.2% Of Reading @ ≥ 5% of Full Stroke Dispense (60mm) +/- 3% Of Reading @ ≥ 1% of Full Stroke Dispense (60mm) | Left Syringe Drive Calibration Result: PASS |
| --- | --- |

Left Syringe Drive Test Data Summary

| Command Instance | Actual 1% Stroke 0.600 mm Command | Actual 5% Stroke 3.00 mm Command | Actual 30% Stroke 18.00 mm Command |
| --- | --- | --- | --- |
| 1 | 0.5985 | 2.9975 | 17.9985 |
| 2 | 0.5985 | 2.9970 | 17.9980 |
| 3 | 0.5985 | 2.9965 | 17.9960 |
| 4 | 0.5990 | 2.9970 | 17.9980 |
| 5 | 0.5990 | 2.9970 | 18.0005 |
| 6 | 0.5995 | 2.9975 | 18.0000 |
| 7 | 0.5990 | 2.9990 | 17.9995 |
| 8 | 0.6000 | 2.9975 | 17.9985 |
| 9 | 0.5995 | 2.9980 | 17.9985 |
| 10 | 0.5995 | 2.9980 | 17.9990 |
| % Allowable Precision / % Actual Precision | 1.5% / 0.082% | 0.5% / 0.022% | 0.2% / 0.007% |
| %Allowable Accuracy / Actual Accuracy | +/- 3% / -0.150% | +/- 1.2% / -0.083% | +/- 1% / -0.007% |
| Result | PASS | PASS | PASS |

Representative Dispensed Volume, 1mL Syringe*

| Description | Actual 1% Stroke 10.000 uL Command | Actual 5% Stroke 50.000 uL Command | Actual 30% Stroke 300.000 uL Command |
| --- | --- | --- | --- |
| Maximum Average Dispense (uL) | 10.044 | 50.255 | 301.761 |
| Minimum Average Dispense (uL) | 9.956 | 49.814 | 299.112 |
| %Allowable Accuracy / %Dispense Accuracy @ Maximum Allowable Syringe Diameter | +/- 3% / 0.442% | +/- 1.2% / 0.508% | +/- 1% / 0.583% |
| %Allowable Accuracy / %Dispense Accuracy @ Minimum Allowable Syringe Diameter | +/- 3% / -0.440% | +/- 1.2% / -0.373% | +/- 1% / -0.297% |

*Representative Dispensed Volume Table is intended to extrapolate measured linear displacement data to the representative dispensed volume from a Hamilton 1mL syringe. Maximum and minimum volumes are calculated from specified syringe barrel inside diameter limits. Accuracy is shown for barrel maximum and minimum inside diameters. Complete data set for each device is on file at Hamilton Company and available upon request.

Electronic Signature By: 2467

Date of Calibration: 01/14/2014

# Text boxes

F-533  REV. A

Label#1: Test Report Envelope

Label#3: Packaging

Label#2: Shipping

SYRINGE PUMP ASSEMBLY ML600 DUAL

PART NO: 61502-01

SERIAL NO.: ML600AK1203	LOT# RMA 49704

*ML600AK1203*

01/14/14

*ML600AK1203*

ML600AK1203

P/N: 61502-01

S/N: ML600AK1203

LOT# RMA 49704

01/14/14